## Slide 1: Paralympics and Sport rehabilitation
26.06.2025
Academic Year 2024-2025

## Slide 2
Amari Giacomo | 10698309
Milazzo Alessandro Giuseppe | 10766415
Traversi Gloria | 10816956
Usai Chiara | 11005835
Zanoni Chiara | 10784594
Introduction
Project proposal
Project design
08.04.2025
25.03.2025
Focus group
Proposal
26.06.2025
Final work
Challenges
Conclusion
Cushion with Independent Air Chambers + Pressure-Sensitive Mat

## Slide 3
Spinal Cord Injuries and FES
Over 15 million people worldwide live with spinal cord injuries (SCI).
Introduction
Project proposal
Project design
Challenges
Conclusion
The trike we use is a modified commercial recumbent tricycle (CATRIKE 700).
Functional electrical stimulation (FES) during pedaling exercises help improve an SCI person’s physical and psychological well-being  FES-cycling.
Sanna, N., Nossa, R., Biffi, E., Guanziroli, E., Diella, E., Ferrante, S., Molteni, F., Peruzzo, D., Nicola, C., Alfonso, M., Rizzo, G., Tarabini, M., Pedrocchi, A., & Ambrosini, E. (2025). Evaluating the Health and Fitness Benefits of a 6-Month FES-Cycling. Program on a Recumbent Trike for Individuals with Motor Complete SCI: A Pilot Study.
Everaert, D. G., Okuma, Y., Abdollah, V., & Ho, C. (2021). Timing and dosage of FES cycling early after acute spinal cord injury: A case series report.
Savona, D., Zanco, C., Sanna, N., Pedrocchi, A., & Ambrosini, E. (2024). A functional Electrical Stimulation and Motor-Assisted trike for sport rehabilitation therapy. 2022 9th IEEE RAS/EMBS International Conference for Biomedical Robotics and Biomechatronics

## Slide 4
User-center need
Risk of PRESSURE ULCERS: prolonged compression that damages the skin and underlying tissues.
Project proposal
Project design
Challenges
Conclusion
Introduction
MITIGATE DISCOMFORT
  Healthy individual: In-Chair Movements (ICM)
  Individual with SCI: lifting themselves from the chair using their arms  not possible during training or competition
The project aims to address this issue by improving safety and comfort of athletes on the trike.
Fenety, A., & Walker, J. M. (2002). Short-Term effects of workstation exercises on musculoskeletal discomfort and postural changes in seated video display unit workers.
Vergara, M., & Page, Á. (2002). Relationship between comfort and back posture and mobility in sitting-posture.

## Slide 5
Requirements
prevent pressure sores
Project proposal
Project design
Challenges
Conclusion
Introduction
The FUNCTIONAL requirements are related to the behaviour and performance of the system and are essential for the correct use of the device.
ensure user comfort
maintain postural stability
be adaptable to the trike
be waterproof and resistant

## Slide 6
Requirements
not bulky and lightweight
Project proposal
Project design
Challenges
Conclusion
Introduction
easy to carry
easy to use
simple to maintain
The NON - FUNCTIONAL requirements refer to aspects that do not directly affect how the system works. They are important to ensure a good user experience.

## Slide 7
Project design
Conclusion
Introduction
Project proposal
Challenges
Cushion with Independent Air Chambers + Pressure-Sensitive Mat
08.04.2025
25.03.2025
Focus group
Proposal
26.06.2025

## Slide 8
The independent cell cushion is designed to adjust the air within the chambers in response to pressure changes detected by the MAT, thereby reducing pressure overload and improving load distribution.
Project design
Conclusion
Introduction
Project proposal
Challenges
Cushion with Independent Air Chambers + Pressure-Sensitive Mat
08.04.2025
25.03.2025
Focus group
Proposal
26.06.2025

## Slide 9
Project design
Conclusion
Introduction
Project proposal
Challenges
Cushion with Independent Air Chambers + Pressure-Sensitive Mat
Material
Sensors
Embedded Systems
Cells
Material
Compressor
ON/OFF Solenoid Valve
Check Valve
Electropneumatic Valve
Mechanical Release Valve
Arduino
Software
1
2
3
4
5

## Slide 10
Cover
The cover is the direct interface with the user.
Project design
Conclusion
Introduction
Project proposal
Challenges
Vergara, M., & Page, Á. (2002). Relationship between comfort and back posture and mobility in sitting-posture.
ITTTAI Technologies - ITTTAI. Available at: https://itttai.com/en/technologies/itttai-technologies
10% binder (Tubicoat PUT), 2% thickener (Tubicoat ASD 60), 2% crosslinker (Fixerer MF 200), 3% softener (Microsil 500), and 4% water repellent (Tubicoat A711).
comfort
breathability
waterproofness
mechanical resistance
hygiene
MICROPOROUS POLYURETHANE (PU) (150–250 µm thickness)

## Slide 11
Cushion
Project design
Conclusion
Introduction
Project proposal
Challenges
Ensuring the stability of the cushion
fastening system,
velcro strips.
Cushion on the shoulder area
Cells on the backrest
Cells on the seat
11
6
1

## Slide 12
Cushion : cells
Flow Advance mattress  SolidWorks  17 air cells in total, 12 of which are dynamically controlled.
Project design
Conclusion
Introduction
Project proposal
Challenges
Robinson, J., Shewalkar, V., Rigo, I., Rock, A., Cinnamon, L., Chapman-Rienstra, D., Hong, J., Kim, J., & Golecki, H. (2023). DESIGN OF a CUSTOM SENSING AND ACTUATING CUSHION FOR USE IN PRESSURE RELIEF IN WHEELCHAIR USERS. 2017
Cushion on the shoulder area:
one without air cells.
Cells on the backrest:
two lateral semicylindrical, not dynamically controlled;
four central semicylindrical, dynamically controlled.
Cells on the seat:
two lateral semicylindrical not dynamically controlled;
one rectangular static;
six rear elliptical dynamically controlled;
two anterior circular dynamically controlled.
1
11
6

## Slide 13
Cushion: material
Project design
Conclusion
Introduction
Project proposal
Challenges
Flow advance : Sistema Antidecubito dinamico (2023) Service Med. Available at: https://www.servicemed.it/prodotti/sistemi-antidecubito-attivi-ad-aria/flow-advance
Sartorato, G. (2022) Sistema antidecubito AD Altissimo Rischio flow advance, Nurse24.it. Available at: https://www.nurse24.it/dossier/wound-care/sistema-antidecubito-flow-advance-ad-altissim
Euro Ausili S.r.l. (n.d.). Genius: Low continuous pressure antidecubitus system with 3 differentiated zones [Brochure]. Available at https://www.euroausili.it/
Based on the description of the Flow Advance system.
elastic
resistant
biocompatible
welding techniques
POLYURETHANE FIBER (PU) (0.2 mm and 0.5 mm thickness)
Each cell has an integrated fitting that connects the chamber to the pneumatic tubing.

## Slide 14
Pressure sensitive MAT: sensors
Two sensor mats: Tekscan CONFORMat™ system in the Research configuration. NOT use all 2000 sensors  actively select only 672 SENSORS grouped into 12 logical areas.
Project design
Conclusion
Introduction
Project proposal
Challenges
https://www.tekscan.com/products-solutions/systems/body-pressure-measurement-system-bpms-research
On the backrest: each cell was associated with 100 sensors.
A total of 400 sensors distributed across 4 cells.
On the seat:
the 6 rear cells are monitored by 32 sensors each;
the 2 larger front cells are monitored by 40 sensors each.
In total, the seat uses 272 sensors.

## Slide 15
Pressure sensitive MAT: embedded systems
Data exported in real-time in ASCII format using a proprietary format (FSX) not compatible with embedded systems like Arduino.
Project design
Conclusion
Introduction
Project proposal
Challenges
Software bridge that converts the data into a format readable by Arduino.
Software Development Kit (SDK) provided by Tekscan and its related Application Programming Interfaces (APIs): enables the development of custom applications in a Windows environment.

## Slide 16
Project design
Conclusion
Introduction
Project proposal
Challenges
Distributed electropneumatic system
COMPRESSOR, the main source of compressed air.
24 V

## Slide 17
Project design
Conclusion
Introduction
Project proposal
Challenges
Distributed electropneumatic system
ON/OFF SOLENOID VALVE, normally closed, which acts as the main pneumatic switch for the entire system.
COMPRESSOR, the main source of compressed air.

## Slide 18
Project design
Conclusion
Introduction
Project proposal
Challenges
Distributed electropneumatic system
ON/OFF SOLENOID VALVE, normally closed, which acts as the main pneumatic switch for the entire system.
CHECK VALVE, allows air to flow only toward the cell.
3/2 ELECTRO PNEUMATIC VALVE, normally closed, ensures air direction.
COMPRESSOR, the main source of compressed air.

## Slide 19
Project design
Conclusion
Introduction
Project proposal
Challenges
Distributed electropneumatic system
ON/OFF SOLENOID VALVE, normally closed, which acts as the main pneumatic switch for the entire system.
CHECK VALVE, allows air to flow only toward the cell.
3/2 ELECTRO PNEUMATIC VALVE, normally closed, ensures air direction.
MECHANICAL RELEASE VALVE keeps the internal pressure within safe limits.
COMPRESSOR, the main source of compressed air.

## Slide 20
Valves and compressor require higher currents than those the microcontroller can supply directly  RELAY CIRCUITS.
Arduino Mega 2560: the central control unit of the entire system.
Project design
Conclusion
Introduction
Project proposal
Challenges
Microcontroller: Arduino
https://store.arduino.cc/products/arduino-mega-2560-rev3?srsltid=AfmBOoriXWEOzhxUq7rx2pqrkn9Ff-agWQZqJYAiDKFTW4Y360G7kTXg
Acquisition of pressure data from the MAT via a USB connection. Processing frequency: 10 Hz.
The microcontroller manages the control of the electro-pneumatic system, using digital output pins :
twelve for the individual control of 3/2 electro-pneumatic valves,
one dedicated pin for the main ON/OFF valve,
one additional pin for the compressor.

## Slide 21
Project design
Conclusion
Introduction
Project proposal
Challenges
Microcontroller: software
THE SYSTEM
During the setup phase, it was ESTIMATED that the cell volume was 5 L at ambient pressure. To reach 0.2 MPa (~10 L at pressure), 17.14 s of inflation are needed using a 35 L/min compressor.
Developed in C++ using the Arduino IDE
The real-time control of the 12 pneumatic cells by processing pressure data at 10 Hz

## Slide 22
Project design
Conclusion
Introduction
Project proposal
Challenges
Microcontroller: software
DURING NORMAL USE
Pressure control on each individual sensor:
calculation and monitoring of average pressures based on the Reswick and Rogers curve,
triggering automatic interventions if the 300 mmHg hour threshold is exceeded.
AT SYSTEM STARTUP
  all cells are inflated for 20 s to ensure operational pressure
the system performs a brief general inflation every 3 min
Gefen, A. (2009). Reswick and Rogers pressure-time curve for pressure ulcer risk.

## Slide 23
Shape and number of air chambers: balance ergonomics and technical simplicity.
Wiring: holes in the trike frame to pass the air tubes under the frame. A net under the frame could hold the tubes and valves neatly, making the system lighter, safer, and better organized.
Software interface: to make software bridge usingTekscan SDK.
Outdoor use and total weight: we need to make lighter versions.
Steps for future improvements
Project design
Challenges
Conclusion
Introduction
Project proposal

## Slide 24
Project proposal
Project design
Challenges
Conclusion
Introduction
08.04.2025
25.03.2025
Focus group
Proposal
26.06.2025

## Slide 25
CUSHION
Cells, Material
PRESSURE SENSITIVE MAT
Sensors, Embedded Systems
MICROCONTROLLER
Arduino, Software
Project proposal
Project design
Challenges
Conclusion
Introduction
08.04.2025
25.03.2025
Focus group
Proposal
26.06.2025

## Slide 26: Paralympics and Sport rehabilitation
26.06.2025
Academic Year 2024-2025

## Slide 27
Final CAD dimensions
Project design
Conclusion
Introduction
Project proposal
Challenges

## Slide 28
Final CAD dimensions
Project design
Conclusion
Introduction
Project proposal
Challenges

## Slide 29
Final CAD dimensions
Project design
Conclusion
Introduction
Project proposal
Challenges

## Slide 30
Alternative CAD designs
Project design
Conclusion
Introduction
Project proposal
Challenges

## Slide 31
Reswick and Rogers curve
Project design
Conclusion
Introduction
Project proposal
Challenges

## Slide 32
Overall cost estimate
Project design
Conclusion
Introduction
Project proposal
Challenges
COMPONENT | COMMERCIAL EXAMPLE | APPROXIMATE PRICE (EUR)
AIR CELL CUSHION | Custom or OEM (e.g., Roho, Vicair) | ~3,000–5,000 €
BREATHABLE COVER | Micro-perforated PU cover | ~200–300 €
MATS | Tekscan CONFORMat system | ~13,000 €
COMPRESSOR (24 V DC) | CAFELE 24 V air compressor | ~38 €
ON/OFF VALVE (2/2) | Heschen 2V025-08 24 V DC | ~12 €
3/2 VALVES (12 UNITS) | Heschen 3V210-08 24 V DC | ~12 € each
CHECK VALVES (12 UNITS) | EFUTURETIME 10 mm quick-connect | ~5 € each
SAFETY VALVES (12 UNITS) | Aurorad 2 bar ½″ | ~6 € each
RELAY BOARDS (2 UNITS) | ARCELI 8-channel 5 V | ~5 €
COMPRESSOR RELAY | CGEAMDY 24 V 30 A automotive relay | ~12 €
WIRING AND PROTECTION | 16 AWG cables, fuses, 1N4007 diodes | ~10–15 € total

## Slide 33
Focus group
ANDREA
Experience: various races
Issues: safety, comfort, upper limbs, outdoor
Purpose: use with families and friends
DAVIDE
Experience: various races
Issues: transporting the trike, customization, comfort, safety, outdoor
Purpose: autonomous training
ALESSANDRO
Experience: training
Issues: safety, comfort, upper limbs, outdoor
Purpose: total body training
Project proposal
Project design
Challenges
Conclusion
Introduction